## Slide 1: ENGINEER CONSTRUCTION EQUIPMENT TEST, MEASURING, AND DIAGNOSTIC EQUIPMENT

## Slide 2
INTRODUCTION
The Army's mission “to perform prompt and sustained combat on land" implies
the readiness to go at any time with what we have in manpower and material
resources. Effective firepower and mobility depend on our ability to
perform the maintenance necessary to keep equipment and material in
operating condition. To maintain combat effectiveness, we need to
strengthen our entire maintenance system, with emphasis on the user level.
Our equipment continually becomes more complex. This all adds to the
requirement for improved maintenance at all levels. The growing demands for
force readiness mean that we must approach our maintenance responsibility
with added enthusiasm. This subcourse teaches the proper use of tools/
equipment necessary to repair and maintain engineer equipment and thus help
the supervisor manage and maintain combat readiness

## Slide 3
Learning Event 1
IDENTIFY TYPES AND MECHANICS OF MICROMETERS
Micrometers are used for measurements requiring precise accuracy. They are
made in various shapes and sizes, depending on their intended purpose. They
all have a precision screw adjustment offering great measuring accuracy.

## Slide 4
TYPES
Micrometer Caliper (Outside)
The micrometer caliper, Figure 1, is the most common micrometer. It has a
range of 0 to 1 inch and is graduated to read in thousandths of an inch or
in units of the metric system, from 0 to 25 millimeters by hundredths of a
millimeter. It may or may not havea.
A stainless steel frame to resist corrosion or tarnish.
b. A ratchet for applying a constant measuring pressure.
c. A special vernier scale for reading tenths of thousandths of an
inch.
d. A clamp ring or locknut for clamping the spindle to hold a setting.
e. Cemented carbide tips on the measuring anvils to reduce w

## Slide 5
The frame can be smaller to the extent that the range of the caliper is only
0 to 1/2 inch; or, it can be larger so that the range is 23 to 24 inches.
The head has a constant range of 0 to 1 inch. The shape of the frame may be
varied to adapt it to the physical requirements of some types of work. For
examplea.
The frame back of the anvil may be cut away to allow the anvil to be
placed in a narrow slot.
b. The frame may have a deep throat to allow it to reach into the
center sections of a sheet (sheet metal or paper gage).
c. The frame may be in the form of a base so that the gage can be used
as a bench micrometer

## Slide 6
d. The frame may have a wooden handle and be of extraheavy
construction for use in gaging hot sheet metal in a steel mill.

## Slide 7
The design of the spindle and anvil may vary to accommodate special physical
requirements. For examplea.
The spindle and anvil may be chamfered so the gage can slide on and
off the work easily, as when gaging hot metal.
b. A ballshaped
anvil is convenient in measuring the thickness of a
pipe section of small diameter.
c. The Vshaped
anvil is necessary on the screw thread micrometer
caliper to mesh properly with the screw thread. The spindle of the
screw thread micrometer is coneshaped.
This micrometer measures
the pitch diameter.
d. Interchangeable anvils of various lengths make it possible to reduce
the range of the caliper. A micrometer having a range from 5 to 6
inches can be changed to one having a 4to
5inch,
or 3to
4inch
range by inserting a special anvil of the proper length.

## Slide 9
Lesson 1/Learning Event 1
Inside Micrometers
Inside micrometers, Figure 2, are used to measure inside dimensions. The
minimum dimension that can be measured is determined by the length of the
unit with its shortest anvil in place and the screw set to zero. It
consists of an ordinary micrometer head, except that the outer end of the
sleeve carries a contact point, attached to a measuring rod.

## Slide 11
The average inside micrometer set, Figure 3, has a range that extends from 2
to 10 inches. The various steps in covering this range are obtained by
means of extension rods. The micrometer set may also contain a collar for
splitting the inch step between two rods. The collar, which is 1/2 inch
long, extends the rod another 1/2 inch so the range of each step can be made
to overlap the next. The range of the micrometer screw itself is very short
when compared to its measuring range. The smallest models have a 1/4inch
screw, and the largest inside micrometers have only a 1inch
screw.

## Slide 13
Depth Micrometers
A depth micrometer, as the name implies, is designed to measure the depth of
holes, recesses, keyways, and so forth. The tool consists of a base with a
micrometer head, Figure 4. Various length measuring rods can be inserted
through a hole in the micrometer screw. The rod protrudes through the base
and moves as the thimble is rotated.

## Slide 15
MECHANICS
Design
The micrometer, Figure 5, makes use of the relation of the circular movement
of a screw to its axial movement. The amount of axial movement of a screw
per revolution depends on the thread, and is known as the lead. For
example, a circular nut on a screw has its circumference divided into 25
equal spaces. If the nut advances axially 1/40 inch for each revolution and
is then turned one division or 1/25 of a revolution, it will move axially
1/25 x 1/40 or 1/1,000 of an inch. In the micrometer, the nut is stationary
and the screw moves forward axially a distance proportional to the amount it
is turned. The screw on a micrometer has 40 threads to the inch, and the
thimble has its circumference divided into 25 parts, so one division on the
thimble represents an advancement of 1/1,000 of an inch axially.

## Slide 17
Construction
Refer to Figure 5. You can see the various components that make up a
micrometer. The following is an explanation of how the components are
assembled to make it work:
a. One end of the Ushaped
steel frame holds a stationary anvil. The
anvil is a hardened button either pressed or screwed onto the frame.
b. The unthreaded part of the screw is the steel spindle. The spindle
advances or retracts to open or close the open side of the Uframe.
The spindle bearing is a plain bearing and is part of the frame.

## Slide 18
c. A hollow barrel extends from this bearing. There is a micrometer
scale on the side of the barrel which is graduated in tenths of an
inch. These are in turn divided into subdivisions of 0.025 inch.
The end of the barrel supports the nut which engages the screw
(adjustment nut). This unit is slotted, and its outer surface has a
tapered thread and a nut. These make it possible to adjust, within
limits, the diameter of the slotted nut to compensate for wear

## Slide 19
d. A thimble is attached to the screw. The thimble is a sleeve that
fits over the barrel. The front edge of the thimble carries a scale
broken down into 25 parts. This scale indicates parts of a
revolution, while the scale on the barrel indicates the number of
revolutions. The thimble is connected to the adjustment screw
through a sleeve that permits it to be slipped in relation to the
screw to allow for adjustment. The inner sleeve is sweated to the
screw. The outer sleeve is clamped to the inner one by the thimble
cap. Loosening the cap makes it possible to slip one in relation to
the other.

## Slide 20
e. Sometimes there is a ratchet on top of the thimble cap. This device
consists of an overriding clutch held together by a spring in such a
way that when the spindle is brought up against the work, the clutch
will slip when the correct measuring pressure is reached. The
ratchet eliminates the effects of differences in personal touch and
thus reduces the possibility of error due to differences in
measuring pressure. Not all micrometers have ratchets.
f. Some micrometers have a clamp ring or locknut located in the center
of the spindle bearing. This clamping device makes it possible to
preserve a setting by locking the spindle in any position desired

## Slide 21
Care
The following rules should be followed to keep micrometers in good condition
and to preserve their accuracy:
a. Never store a micrometer with its anvil and spindle closed. Leave a
small gap between the anvil and spindle when storing a micrometer
since flat surfaces which have been pressed together for any length
of time tend to corrode.

## Slide 22
b. Oil the micrometer in only one placethe
screw. If storing a
micrometer for long periods of time, however, cover it with a light
film of oil and wrap it in oiled paper.
c. Never roll the thimble along the hand or arm. Likewise, never hold
the thimble and twirl the frame to open or close the micrometer.
This misuse causes excessive wear on the screw.
d. Before using a micrometer, wipe it off and pull a piece of paper
between the anvil and the end of the spindle.

## Slide 23
e. The micrometer should operate freely with no play in its travel. If
a micrometer has play, or if it binds, it should be returned to the
manufacturer for reconditioning. (Play is caused by abuse or uneven
wear.)
f. Check the micrometer screw periodically with a precision gage block
in at least four places other than zero to verify its accuracy.
Simply measure a selected group of blocks ranging from 0 to 1 inch.
g. Clean the micrometer mechanism whenever it becomes gummy, contains
abrasive grit, or needs to be adjusted. Use an approved cleaning
agent.

## Slide 24
h. Test the accuracy of the micrometer periodically. When the faces of
the spindle and anvil become worn and are no longer flat and
parallel to each other, measurement errors can result. The error
tolerance must not exceed 0.0002 inch on a micrometer graduated to
control measurements to a limit of 0.001 inch and must not exceed
0.00005 inch in a micrometer which is graduated to control
measurements to a limit of 0.0001 inch. Measuring a ball at several
points over the surface of the anvil will reveal any error in
parallelism. Parallelism can be tested using two balls mounted in
an aluminum holder. If the anvils are in error more than the
allowable maximum, the micrometer should be returned to the
manufacturer for repair.

## Slide 25
i. Adjust the micrometer thimble setting periodically. In adjusting a
micrometer to read correctly, the thimble is not set to 0 when the
anvil is in contact with the spindle, but is set at some dimension
to distribute the error. For example, if a micrometer screw has an
accumulating error of 0.0003 inch in

## Slide 26
the length of its travel, and it were set correctly at 0, it would
be off 0.0003 at 1 inch. However, if the micrometer were set
correctly in the center of its travel, it would be 0.00015 under at
0 and 0.00015 over at 1 inch, which is a much better condition.
Micrometers do not always return exactly to 0 when the anvil and
spindle contact. This does not mean that it is not adjusted
properly. Turn the friction sleeve with a small spanner wrench to
compensate for minor wear on the anvil and spindle or on the screw,
Figure 6

## Slide 28
Learning Event 2
MEASURE USING THE MICROMETER
SELECTION
Commonly used micrometers are made so that the longest movement possible
between the spindle and the anvil is 1 inch. This movement is called the
range. Micrometer frames, however, are available in a wide variety of
sizes, from 1 inch up to as large as 24 inches. The range of a 1inch
micrometer is from 0 to 1 inch. In other words, it can be used to measure
work that is 1 inch or less. A 2inch
micrometer has a range from 1 to 2
inches. A 6inch
micrometer has a range from 5 to 6 inches, and will only
measure work between 5 and 6 inches.

## Slide 29
Mechanics must first find the approximate size of the work to the nearest
inch, and then select the correct micrometer. For example, to find the
exact diameter of a piece of round stock use a rule and find the approximate
diameter of the stock. If it is between 3 and 4 inches, a micrometer with a
3to
4inch
gage would be required to measure the exact diameter.
Similarly, when using inside and depth micrometers, rods of suitable lengths
must be fitted into the tool to get the approximate inside dimension within
an inch. After obtaining the approximate dimensions, an exact measurement
is read by turning the thimble. The size of the micrometer indicates the
size of the largest work it will measure.

## Slide 30
READING STANDARD MICROMETERS
Reading a micrometer consists of simply reading the micrometer scale or
counting the revolutions of the thimble and adding to this any fraction of a
revolution. The micrometer screw has 40 threads per inch. This means that
one complete revolution of the micrometer screw moves the spindle away from
or towards the anvil exactly 0.025 inch. The lines on the barrel, Figure 7,
conform to the pitch of the micrometer screw with each line being numbered.
The bevel edge of the thimble is graduated into 25 parts, each part
indicating 0.001 inch or 1/25 of the 0.025 inch covered by a complete
revolution. Every fifth line on the thimble is numbered to read a
measurement in thousandths of an inch.

## Slide 32
Use the following sequence to read the measurement shown in Figure 7.
a. Highest figure visible on barrel. 2 = 0.200 inch
b. Number of lines visible between
the No. 2 and thimble edge. 1 = 0.025 inch
c. Line on the thimble that coincides
with or has passed the revolution
or long line on the barrel. 16 = 0.016 inch
d. Measurement reading Total = 0.241 inch

## Slide 33
READING METRIC MICROMETERS
The same principal applies to reading metric graduated micrometers. But,
the following changes in graduation should be noted to avoid confusion:
a. The pitch of the metric micrometer screw is 0.5 millimeter (mm).
One revolution of the spindle advances or withdraws the screw a
distance equal to 0.5 mm.
b. The barrel is graduated in millimeters, from 0 to 25. It takes two
revolutions of the spindle to move 1 mm.

## Slide 34
c. The thimble is graduated in 50 divisions with every fifth line being
numbered.
d. Rotating the thimble from one graduation to the next moves the
spindle 1/50 of 0.5 mm, or 1/100 mm. Two graduations equals 2/100
mm and so forth.

## Slide 35
Refer to Figure 8, and follow the same sequence as was used for the standard
micrometer and determine the measurement indicated.
a. Highest figure visible on barrel. 20 = 20 mm
b. Number of lines visible between
No. 20 and thimble edge. 2 = 2 mm
c. Line on thimble that coincides
with or has passed the revolution
or long line on the barrel. 36 = .36 mm
d. Measurement reading. Total = 22.36 mm

## Slide 37
READING VERNIER MICROMETERS
Reading vernier micrometers is the same as reading standard micrometers. An
additional step must be taken to add the vernier reading to the dimensions.
This allows precise measurements accurate to ten thousandths (0.0001) of an
inch. This scale furnishes the fine readings between the lines on the
thimble rather than making an estimate as you would on a standard
micrometer.
The 10 spaces on the vernier scale, Figure 9, (1) are equivalent to 9 spaces
on the thimble (2). Therefore, each unit on the vernier scale is equal to
0.0009 inch and the difference between the sizes of the units on each scale
is 0.0001 inch.

## Slide 39
Refer to Figure 10. Read the measurement shown on the vernier micrometer.
a. Highest figure visible on barrel (3). 2 = 0.200 inch
b. Number of lines visible between
No. 2 and thimble edge (4). 3 = 0.075 inch
c. Line on the thimble that coincides
with or is nearest the revolution
or long line in the barrel (5). 11 = 0.011 inch
d. Line on the vernier scale that
coincides with the line on the
thimble (6). 2 = 0.0002 inch
e. Add all readings together and total. .2862 inc

## Slide 41
Additional micrometer reading exercises later in this lesson will assist you
in understanding how to read micrometers.
ADJUSTING THE MICROMETER CALIPER TO WORK
Figure 11, Part A, shows the proper way to hold a micrometer caliper in
checking a small part. Hold the part in one hand. Hold the micrometer in
the other hand so that the thimble rests between the thumb and forefinger.
The third finger is then in a position to hold the frame against the palm of
the hand. When the frame is supported this way it is easy to guide the work
over the anvil. The thumb and forefinger are thus in position to turn the
thimble either directly or through the ratchet and bring the spindle over
against the work.

## Slide 43
On larger work, it is often necessary to have the work stationary and
positioned to permit access to the micrometer. Figure 11, Part B, shows the
proper method of holding a micrometer when checking a part too large to be
held in one hand. The frame is held by one hand to position it and to
locate it square to the measured surface. The other hand operates the
thimble either directly or through the ratchet. You should check large flat
parts in several places to determine the amount of variation.
To gage a shaft, Figure 11, Part C, hold the frame in one hand while
operating the thimble with the other. In measuring a cylindrical part with
a micrometer, it is necessary to “feel" the setting to be sure that the
spindle is on the diameter. You must check the diameter in several places
to determine the amount of outofroundness.

## Slide 44
Micrometer calipers are made in various sizes up to 168 inches. Figure 12
shows a pulley being checked with a micrometer whose range has been reduced
by a special anvil screwed into the frame. A set of different length anvils
permits this micrometer to be used over a wide range of sizes. Yet, the
spindle only moves 1 inch. This micrometer has been lightened in weight by
its Isection
construction and by boring holes in the frame.

## Slide 46
USING INSIDE MICROMETERS
The normal procedure for using inside micrometers is to set them across
diameters or between inside surfaces, remove them, and then read the
dimension. For this reason, the thimble on an inside micrometer is much
stiffer than the one on a micrometer caliper. Thus, it holds the dimension
better. It is good practice to verify the reading of an inside micrometer
by measuring it with a micrometer caliper.

## Slide 47
Figure 13 shows an inside micrometer with extension rod being used to check
the diameter of a bored hole. Note the arrows which indicate the direction
the operator is measuring (feeling) for the largest dimension horizontally
and the smallest dimension vertically. Inside micrometers have spherical
contact points which require more practice to "feel” the full diametral
measurement. One contact point is generally held in a fixed position and
the other rocked in different directions. This ensures the tool spans the
true diameter of a hole or the correct width of a slot.

## Slide 49
A handle can be clamped onto the body of the micrometer when probing a deep
hole or in a restricted place.
Use the following procedures to transfer measurements from inside calipers
to micrometer calipers:
a. After setting the inside caliper or inside micrometer to the work,
hold the micrometer caliper in one hand and the inside tool in the
other hand.
b. Turn the thimble of the micrometer caliper with the thumb and
forefinger until you feel the inside tool legs lightly contact the
anvil and spindle of the micrometer caliper.
c. Hold the tips of the inside tool legs parallel to the axis of the
micrometer caliper.
d. The micrometer caliper will be accurately set when the inside tool
will just pass between the anvil and spindle by its own weight

## Slide 50
Lesson 2: CALIPERS AND DIVIDERS

## Slide 51
Learning Event 1
IDENTIFY VARIOUS TYPES OF CALIPERS AND DIVIDERS
Calipers are devices for determining the dimensions of machine parts.
Calipers are not usually precision instruments but are sufficiently accurate
for the fabrication of many parts. Simple calipers consist of two movable
legs joined at one point with the separation between them controlled by a
springloaded
screw.

## Slide 52
The ends of the legs of inside calipers turn outward to measure widths of
grooves, interior diameters, and similar dimensions. Outside calipers have
curved legs with turnedin
ends suited to measuring outside diameters.
Scribes are calipers with straight legs and sharp points used for marking
off distances on flat surfaces. Hermaphrodite (hur maf re dit) calipers
have one straight scribing leg. They are used to scribe a line parallel to
a curved shoulder. In simple calipers, the opening must be placed against a
scale to read the dimension.
Vernier calipers incorporate a scale plus a vernier and are capable of
precise measurement. They are shaped something like an adjustable wrench
with points at the end of the jaws. These points are suited to both inside
and outside surfaces

## Slide 53
CALIPERS
Outside
Outside calipers, Figure 18, are used to measure distances over and around
adjacent outside surfaces. They are also used to transfer the measurements
to a rule. Several types and sizes are made to accommodate a wide range of
measurements. The size of the caliper is expressed in terms of the maximum
dimension it is capable of measuring. A 3inch
caliper, for example, will
measure a distance of 3 inches. Actually, the maximum capacity of the
caliper will be greater, often as much as onethird.
This means, for
example, that a 3inch
caliper will actually measure up to 4 inches. It is
not recommended that you use a 3inch
caliper to measure 4 inches, since
that could spring the legs resulting in an inaccurate measurement

## Slide 55
Inside
Inside calipers, Figure 19, are used to measure distances between inside
surfaces. The points are round and slightly ballshaped.
This ball shape
establishes the point of contact. Where surfaces are likely to have an
inside curvature, error can occur if the radius of the hole being calipered
is less than the radius of the ballshaped
points. Some inside calipers are
equipped with an adjusting screw on one leg to provide fine adjustment of
the caliper legs.

## Slide 57
Lesson 2/Learning Event 1
Inside and outside calipers have the following design characteristics:
Spring
a. Spring calipers, Figures 18 and 19, are available in sizes from 2 to
8 inches. The friction of the adjusting nut and screw works against
the tension of the spring which holds the legs in any set position.
This type of caliper is known as the toolmaker's spring caliper.
b. Thread spring calipers are used to measure diameters and distances
in tapped holes. The ends of the legs of thread calipers are shaped
to a fine point so that exact contact can be made between threads

## Slide 58
Firm Joint
Firm joint calipers are available in a number of sizes from 3 to 24 inches.
This type of caliper is equipped with a nut and stud that provides
sufficient friction to hold the legs in any set position. Some of these
calipers are equipped with an adjusting screw for fine adjustments

## Slide 59
Transfer Firm Joint
Transfer firm joint calipers are shaped for inside measurements and are used
for measuring recesses where the setting cannot be transferred to a scale
directly because the legs must be collapsed to remove them from the work.

## Slide 60
Hermaphrodite Calipers
Hermaphrodite calipers, Figure 20, are a cross between a divider and an
inside caliper. They have one leg of each device. These calipers are used
for scribing parallel lines from an edge or for locating the center of
cylindrical work. Some are equipped with an adjustable point

## Slide 62
Slide
Slide calipers, Figure 21, have one fixed jaw and one sliding jaw. When the
two jaws are brought in contact with surfaces to be measured, the distance
between the surfaces may be read from the scale. The ends of the jaws are
shaped so that it is possible to measure both inside and outside surfaces.
Slide calipers are normally made in 3inch
sizes, although larger sizes are
available. The standard 3inch
slide caliper measures from 0 to 20 inches
outside, and from 1/8 inch to a little more than 2 inches inside. The
caliper shown has a mark or graduation line on the fixed jaw. This enables
the user to read the inside measurement, while the outside measurement can
be read directly from the inner edge of the fixed jaw. If these marks were
not on the slide caliper, it would be necessary to add the thickness of the
nibs to the reading when using it as an inside caliper. All slide calipers
are equipped with a locking device which makes it possible to hold the jaws
in any desired position. There is an ordinary rule on the other side.

## Slide 64
Vernier
This type of caliper uses a vernier scale. This scale, Figure 22, consists
of a short auxiliary scale usually having one more graduation in the same
length as the longer main scale.
The vernier caliper consists of a Lshaped
frame, the end of which is a
fixed jaw. The long arm of the L is inscribed with the main true scale or
fixed scale. The sliding jaw carries the vernier scale on either side. The
scale on the front side is for outside measurements; the one on the back is
for inside measurements. On some vernier calipers, the metric system of
measurement is placed on the back side of the caliper in lieu of a scale
used for inside measurements. In such cases, add the thickness of the nibs
to the reading when making inside measurements.

## Slide 65
The tips of the jaws are formed so that inside measurements can be taken.
The inside scale automatically compensates for the thickness of the
measuring points.
The length of the jaws range from 1 1/4 inches to 3 1/2 inches. The minimum
inside measurement, using the smallest caliper, is 1/4 inch or 6
millimeters.
Vernier calipers are made in standard sizes of 6, 12, 24, 36, and 48 inches,
and 150, 300, 600, and 900 millimeters. The jaws of all vernier calipers,
except the larger sizes, have two center points, which are particularly
useful in setting dividers to exact dimensions.

## Slide 67
DIVIDERS
Dividers are used for measuring distances between two points, for
transferring or comparing measurements directly from a rule, or for scribing
an arc, radius, or circle. Two types of dividers are spring and wing and
are described below.

## Slide 68
Spring
A spring divider, Figure 23, consists of two sharp points at the end of two
straight legs, held apart by a spring and adjusted using a screw and nut.
The spring divider is available in lengths from 3 to 10 inches.

## Slide 70
Wing
Wing dividers, Figure 24, have a steel bar separating their legs, a locking
nut for securing a rough measurement, and an adjusting screw for fine
adjustments. Wingtype
dividers are available in 6,
8,
and 12inch
lengths. An improved version of this type divider has one leg with a
removable tip so a pencil can be inserted

## Slide 72
Care of Dividers and Calipers
a. Never use a divider or caliper for anything but its intended
purpose.
b. Never pile these tools in a drawer.
c. Never force dividers and calipers beyond their capacity or setting.
d. Never use these tools incorrectly. For example, changing settings
by hammering instead of loosening a clamping screw or nut, bearing
down too hard when scribing with a divider, wearing measuring
surfaces unnecessarily by using heavy measuring pressure.
e. Apply a protective film of oil to tools, when not using them

## Slide 73
Care of Vernier Calipers
The accuracy of the vernier caliper depends on the fit of the sliding jaw
and the wear and distortion in the measuring surfaces. The sliding jaw
should move easily and still not have play. It may be adjusted by removing
the gib in the sliding jaw assembly and bending it. The gib holds the
adjustable jaw against the inside surface of the blade with just the right
pressure to give it the proper friction.

## Slide 74
Wear on the jaws of the vernier caliper is mostly at the tips where most
measurements are made. A certain amount of this wear may be taken up by
adjusting the vernier scale itself. This scale is mounted with screws in
elongated holes. This permits slight adjustments to compensate for wear and
distortion. When the error exceeds 0.0002 inch, either in parallelism or
flatness, the caliper should be returned to the manufacturer for
reconditioning. Wear on the jaws can best be checked visually or by
measuring rolls or rings of known dimensions.

## Slide 75
Learning Event 2
MEASURE USING CALIPERS AND DIVIDERS
DIVIDERS
In setting a divider to a dimension on a scale, the usual procedure is to
locate one point in one of the inch graduations of the rule and to adjust
the nut or screw so that the other point falls easily into another inch
graduation, Figure 25, Part A.
NOTE. Do not set the point to the end of the rule. This could cause an
inaccurate measurement.
Make certain the points of the divider are not blunt, Figure 25, Part

## Slide 77
When transferring a dimension from a part or tool to the scale on a rule,
use the same care in adjusting the points of the dividers. Make sure there
is no excess pressure. Excess pressure tends to spring the points either in
or out. Figure 25, Part C, shows a mechanic scribing a radius on a die
block he is laying out.

## Slide 78
CALIPERS
Outside
These calipers are used in one of two ways. Either the caliper is set to
the dimension of the work and the dimension transferred to a scale, or the
caliper is set on a scale and the work machined until it checks with the
dimension set on the caliper. To adjust an outside caliper to a scale
dimension, one leg of the caliper should be held firmly against one end of
the scale and the other leg adjusted to the desired dimension, Figure 26,
Part A.

## Slide 80
To adjust the outside caliper to the work, open the legs wider than the work
and then bring them down to the work. A sense of feel must be acquired to
use calipers properly. This comes through practice and care in using the
tool to eliminate the possibility of error. Always position the caliper
properly on the axis of the work, Figure 26, Part B.
NOTE. Never set a caliper on work that is revolving in a machine. The
contact of one leg of a caliper on a revolving surface will tend to draw the
other leg over the work because of the friction between the moving surfaces.
Only a slight force is necessary to spring the legs of a caliper thus
measurements made on moving surfaces are never accurate.

## Slide 81
Inside
The inside caliper is set to a dimension by placing the end of a scale and
one point of the caliper against a solid surface and adjusting the other leg
to the proper graduation, Figure 27, Part A.

## Slide 83
To adjust the inside caliper to the work, open the legs wider until they
touch the work, Figure 27, Part B. Always position the caliper on the axis
of the work.
Figure 28 illustrates correct and incorrect positioning of the calipers with
relation to the axis of the work.

## Slide 85
The transfer feature of the caliper is illustrated in Figure 29. Note that
the diameter being measured is recessed and the setting cannot be
transferred to a scale directly because the legs must be collapsed to get
them out of the work. The setting must be reproduced after the calipers are
removed.

## Slide 87
Figure 30 shows how a micrometer can be used to transfer a dimension from an
inside spring caliper

## Slide 89
Hermaphrodite
The hermaphrodite caliper is adjusted and set the same as outside and inside
calipers, depending on the position and use of the caliper leg, Figure 31.

## Slide 91
Vernier
Vernier calipers permit precise, accurate readings using a graduated steel
rule and a movable jaw with the vernier scale. In order to use the vernier
caliper, a thorough vernier scale and the ability to read it are essential.
The steel rule of the caliper is graduated in fortieths or 0.025 of an inch.
Every fourth division represents a tenth of an inch, and is numbered as
shown in Figure 32.

## Slide 93
The vernier scale is divided into 25 parts and numbered 0, 5, 10, 15, 20,
and 25. These 25 parts are equal to and occupy the same space as 24 parts
on the rule. The difference between the width of one of the 25 spaces on
the vernier scale and one of the 24 spaces on the rule is 1/25 of 1/40, or
1/1,000 of an inch. If the tool is set so that the 0 line on the vernier
coincides with the 0 line on the rule, the line to the right of 0 on the
vernier scale will differ from the line to the right of 0 on the rule by
1/1,000 of an inch; the second line by 2/1,000 of an inch; and so forth.
The difference will continue to increase 1/1,000 of an inch for each
division until the 25 on the vernier scale coincides with line 24 on the
rule. To read the scales, note how many inches, tenths (or 0.100), and
fortieths (or 0.025) the mark 0 on the vernier scale is from 0 mark on the
rule; then note the number of divisions on the vernier scale from 0 to a
line which exactly coincides with a line on the rule.

## Slide 94
For example, Figure 32, shows the 0 mark of the vernier scale coinciding
with a line on the rule (see arrow). In this case, the vernier scale is not
necessary because there is no fractional part of a space to determine. The
reading is 2.350.
The 0 mark on the vernier scale, Figure 33, coincides with a fractional part
of a space on the rule. The reading is 2.35 plus a fraction of the space on
the rule

## Slide 96
To determine what fractional part of a whole rule division, or how many
thousandths are to be added to the 2.35 reading, it is necessary to find the
line on the vernier scale that exactly coincides with the line on the rule.
In this case the line coincides at the eighteenth mark. This indicates
18/25 of a whole space. Since each space on the rule equals 0.025 inch,
this part of a space is equal to 0.018 inch, and the total reading is 2.35
plus 0.018, or 2.368 inches.
NOTE. Vernier scales are not necessarily 25 divisions long; they may have
any number of units. For example, the ten thousandths micrometer has a
scale of only ten divisions.

## Slide 97
The vernier caliper has a wide range of measurement applications. The shape
of the measuring jaws and their position with respect to the scale makes
this tool more adaptable than a micrometer. The vernier caliper, however,
does not have the accuracy of a micrometer. In any 1 inch of its length a
vernier caliper should be accurate within 0.001 inch. In any 12 inches, it
should be accurate within 0.002 inch and increase about 0.001 for every
additional 12 inches. The accuracy of measurements made with a vernier
caliper is dependent on the user's ability to feel the measurement

## Slide 98
Because the jaws are long and may have some play in them (especially if an
excessive measuring pressure is used), it is necessary to develop the
ability to handle the vernier caliper. This ability (feel) may be acquired
by measuring such known standards as gage blocks and plug gages.
Applications of the vernier caliper are shown in Figure 34. Figure 34, Part
A, shows a machinist checking the outside diameter of a part. One hand
holds the stationary jaw to locate it, while the other hand operates the
adjusting nut and moves the sliding jaw to the work. The same procedure is
used in checking the inside dimension, Figure 34, Part B.

## Slide 101
Reading a Metric Vernier Caliper
Refer to Figure 35, the steel rule (1) is divided into centimeters (cm) (2).
The longest lines each represent 10 millimeters. Each millimeter is divided
into quarters. The vernier scale (3) is divided into 25 parts and is
numbered 0, 5, 10, 15, 20, and 25.

## Slide 103
Again refer to Figure 35. The metric vernier caliper would be read as
follows:
a. The total number of millimeters
(1) to the left of the vernier
zero index (2) and record. 32.00 cm
b. The number of quarters
(3) between the millimeter
mark and the zero index and
record. .25 mm
c. The highest line on
the vernier scale (4) which
lines up with the line on
the scale (5) and record. .18 mm
d. Add the three readings. Total 32.43

## Slide 104
Lesson 3: DIAL INDICATORS

## Slide 105
Learning Event 1
IDENTIFY DIAL INDICATORS BY TYPE AND CONSTRUCTION
The dial indicator, Figure 40, is a precision measuring instrument. It is
used in machine shops to check size variations of machined parts from the
desired size. It is also used to measure the runout of a flywheel, the end
play of a crankshaft, the lift of a camshaft lobe, the alignment accuracy of
work, and the machine in machine setups.
Dial indicators range in diameter from about 1 inch to 4 1/2 inches. Each
important part has been named and is indicated in the illustration (Figure
40).

## Slide 107
Dial indicators are usually classified as either 1/1,000 indicators or
1/10,000 indicators. This is marked on the dial face, just below the hand,
Figure 41. The dial may be of the balance type, Figure 41, in which the
figures read both to the left or to the right of the zero. Or, it may be of
the continuous reading type, Figure 42, in which the figures read from zero
only in the clockwise direction. The dial is marked off (graduated) so that
each graduation (space between to adjacent lines) represents a definite
movement of the plunger, which is designed so that it will slide in and out.
On 1/1,000type
indicators, each graduation may represent a plunger movement
of onethousandth,
onehalf
thousandth, or onequarter
thousandth of an
inch. On 1/10,000 indicators, each graduation may represent a plunger
movement of one tenthousandth,
onehalf
of one tenthousandth,
or as small
as onequarter
of one tenthousandth
of an inch.

## Slide 110
Learning Event 2
MEASURE USING DIAL INDICATORS
The dial indicator can be used for detecting differences in the size of
various parts of a workpiece. For example, assume we have a 6inch
long
shaft of 3/4inch
diameter which we wish to check for concentricity
(roundness) and taper. To check or indicate for concentricity, you must use
the setup shown in Figure 43. The Vblock
clamp screws (not shown) must be
adjusted just enough to permit you to rotate the end of the shaft by hand as
it rests in the blocks. Loosen the spindle lockscrew of the pedestal, so
you can position the indicator at the proper height. Determine the proper
height by watching the pointer of the indicator. Bring the bottom of the
indicator to bear on the workpiece between the Vblocks
until the pointer
moves clockwise .015. At that point, lock the indicator to the spindle of
the pedestal by turning the spindle lockscrew. Now loosen the bezel clamp
and turn the bezel until the 0 on the face of the dial lines up directly
under the pointer of the indicator. After this, clamp the dial in place by
turning the bezel clamp screw.

## Slide 111
To check the shaft for concentricity, simply turn the end of the shaft by
hand and note the amount of variation detected by the deflection or movement
of the pointer as the workpiece is revolved

## Slide 113
To check this same piece for taper, first remove the Vblock
clamps. It is
only necessary to lay the piece in the blocks. Then follow the procedures
outlined above. Adjust the indicator to the proper height on the pedestal
and bring the button of the indicator to bear on one end of the workpiece.
Adjust the dial so the 0 on the dial lines up with the pointer of the
indicator. Now move the indicator pedestal base so the button of the
indicator is brought to bear on the opposite end of the workpiece. By
noting the deflection of the needle from 0, you can readily see the amount
of taper in the shaft.
Always mount the dial indicator with the plunger at right angles to the
object being measured. Failure to do this can result in inaccurate
readings.

## Slide 114
The value of the graduations on the indicator dial of the dial indicator is
easily determined. Using Figure 44 as an example, first look at the figure
on the dial face which identifies the indicator as a 1/1,000 type. (The
figure is just below the center of the hand.) Then count the number of
spaces between the 0 and the first numeral to the left of the 0. The first
numeral is a 10, indicating .010 of an inch (10/1,000) of plunger travel.
Since there are 10 spaces between the 0 and the 10, the value of each
graduation is equal to onethousandth
inch.

## Slide 116
Since it is obvious that such small movements of the plunger cannot be seen,
the plunger is linked to the indicator hand or pointer through a train of
small gears or through a set of linking levers which multiply any small
movement of the plunger into a larger, sore easily seen movement of the
indicator hand. This is why the dial indicator must be mounted rigidly to
some support (magnetic base), Figure 45. Thus only the plunger can move.
Any movement of the indicator body, in relation to the plunger, would cause
an error in the reading.

## Slide 118
When not in use, the plunger spring naturally pushes the plunger outward,
away from the indicator body, to the limit of plunger travel. On most
standard indicators, the indicator hand comes to rest as shown in Figure 40.
From this position, the hand could only travel to the right, or clockwise.
When you use the indicator properly, the pointer could move either to the
right or left of the 0 (to indicate plus or minus). The indicator must be
set up for use so the plunger has been pushed inward far enough to allow
equal travel to the right or to the left of 0. When using the indicator,
remember that as the plunger moves toward the indicator body, the pointer
moves clockwise, and as it moves away from the indicator body, the pointer
moves counterclockwise.

## Slide 119
Handle dial indicators carefully, and store them in their proper cases when
you finish using then.
As stated earlier, the dial indicator is used to make measurements of
flywheel runout, crankshaft end play, and lift of camshaft lobes. They my
also be used to measure runout of disk brake rotors, cylinder bores of an
engine, and so forth.

## Slide 120
Lesson 4: MEASURING GAGES

## Slide 121
Learning Event
IDENTIFY TYPES AND USES OF MEASURING GAGES
In this lesson we will discuss the following measuring gages:
GAGES
Thickness (Feeler)
These gages are fixed in leaf form which allows you to check and measure
small openings such as valve clearances, narrow slots such as point gaps,
and so forth. They are widely used to check the flatness of parts in
straightening and grinding operations and in squaring objects with a try
square.

## Slide 122
Wire and Drill
Wire gages are used for gaging metal wire. A similar gage is also used to
check the size of hot and cold rolled steel, sheet and plate iron, and music
wire. Drill gages determine the size of a drill and indicate the correct
size of drill bit to use for a given tap size. Drill number and decimal
size are also shown on this type gage

## Slide 123
Thread
Among the many gages used in machining and inspecting threads are center
gages and screw pitch gages.
a. The center gage is used to set thread tools. Four scales on the
gage are used for determining the number of threads per inch.
b. Screw pitch gages are used to determine the pitch of an unknown
thread. The pitch of a screw thread is the distance between the
center of one tooth to the center of the next tooth.

## Slide 124
Small Hole Set
This set of four or more gages is used to check dimensions of small holes,
slots, grooves, and so forth. The set can measure diameters from
approximately 1/8 to 1/2 inch.
Telescoping
These gages are used for measuring the inside size of slots or holes up to 6
inches in width or diameter.

## Slide 125
Fillet and Radius
These gages are used to check convex and concave radii in corners or against
shoulders.
Drill Point
These gages are used to check the accuracy of drill cutting edges after
grinding. They are also equipped with a 6inch
hook rule. This tool can be
used as a drill point gage, hook rule, plain rule, and a slide caliper for
taking outside measurements

## Slide 126
TYPES OF GAGES
Thickness (Feeler)
Thickness (feeler) gages are made in many shapes and sizes. Usually 2 to 26
blades are grouped into a single tool and graduated in thousandths of an
inch. Most thickness blades are straight. Others are bent at the end at
45degree
and 90degree
angles. Some thickness gages are grouped so that
there are several short and several long blades together. Thickness gages
are also available in single blades and in strip form for specific
measurements. For convenience, many groups of thickness gages are equipped
with a locking screw in the case. This screw locks the blade to be used in
the extended position. Figure 49 shows several types of thickness gages
available through the Army Supply System.

## Slide 128
a. The twist drill and drill rod gages, Figure 50, Part A, each have a
series of holes with size and decimal equivalents stamped adjacent
to each hole. One gage measures drill sizes numbers 1 to 60; the
other gage measures drill sizes 1/16 to 1/2 inch by 1/64inch
intervals.
b. Wire gages, Figure 50, Part B, are circular in shape with cutouts in
the outer perimeter. Each cutout gages a different size wire, from
0 to 36 of the English standard wire gage. A separate gage is used
for American standard wire and another for US standard sheet and
plate iron and steel

## Slide 130
Thread
a. Center gages, Figure 51, Part A, are graduated in fourteenth,
twentieth, twentyfourth,
and thirtieth of an inch. The back of the
center gage has a table giving the double depth of thread in
thousandths of an inch for each pitch. This information helps you
determine the size of tap drills. Sixtydegree
angles in the shape
of the gage are used for checking Unified and American threads or US
standard threads and for checking thread cutting tools

## Slide 131
b. Screw pitch gages, Figure 51, Part B, are made for checking the
pitch of US standard, metric, national form, Vform,
and whitworth
cut threads. These gages come grouped in a case or handle, like
thickness gages. The number of threads per inch is stamped on each
blade. Some types have blade locks. The triangular shaped gage has
51 blades covering a very wide range of pitches, including 11 1/2
and 27 threads per inch for Vform
threads.

## Slide 133
a. Small hole gages, Figure 52, Part A, are adjustable, having a
rounded measuring member. A knurled screw in the end of the handle
is turned to expand the ballshaped
end in small holes and recesses.
A micrometer caliper is used to measure the ball end. Maximum
measuring capacity is 1/2 inch.
b. Telescoping gages, Figure 52, Part B, are used to gage larger holes
and to measure inside distances. A knurled screw in the end of the
handle locks the plunger in the measuring position. Maximum
measuring capacity is 6 inches. As with small hole gages,
measurements must be calipered on the gage using a micrometer

## Slide 135
Fillet and Radius
The blades of fillet and radius gages are made of hard rolled steel.
Doubleended
blades, Figure 53, Part A, have a lock which holds the blades
in position. Inside and outside measurements are made with the same blade.
The other gage, Figure 53, Part B, has separate blades for inside and
outside measurements. Each blade of each gage is marked in sixtyfourth.
Gage A, Figure 53, has a range of sizes from 17/64 to 1/2 inch. Gage 3,
Figure 53, ranges in size from 1/32 to 17/64 inch. Each gage has 16 blades

## Slide 137
Drill Point
Drill point gages, Figure 54, consist of a 6inch
hook rule with a 59degree
sliding head that slides up and down the rule. The sliding head can be
locked at any position on the rule. The sliding head is graduated in 1/32
inch increments.

## Slide 139
a. Exercise care when using thickness gages to measure clearance of
knives and cutters on machines. Do not lower a knife on a thickness
blade and then try to remove the gage since the blade may be shaved
off if it is too tight.
b. Never use gages for cleaning slots or holes.
c. When blades are damaged or worn, replace them.
d. Blades in cases are removed by loosening the clamp and sliding out
the damaged blade. Install a new blade and tighten the clamp.
e. To prevent rust when storing gages, always coat the metal parts with
a light film of oil

## Slide 140
f. Store gages in separate containers.
g. Do not stack or pile gages.
h. Always return blades of leaftype
gages to their cases after use.
i. Keep graduations and markings on gages clean and legible.
j. Do not drop gages. Minute scratches or nicks will result in
inaccurate measurements

## Slide 141
GAGE USES
Thickness (Feeler)
Thickness (feeler) gages are used in one of two ways, as a means for
determining a measure or as a means for adjusting to a definite limit.
Figure 55, Part A, shows a thickness gage being used to check piston ring
gap clearance in a cylinder bore. A long blade thickness gage is being used
to determine the fit between large mating surfaces in Figure 55, Part B. By
combining blades, it is possible to measure a wide variation of thicknesses

## Slide 143
Wire and Drill
The use of a drill gage is illustrated in Figure 56, Part A. The size of a
drill is being determined. The drill size or number and decimal size are
stamped on the gage beside each hole. A chart on the gage indicates the
correct size of drill to use for a given tap size. To determine the size of
both sheet stock and wire, you can use the wire gage shown in Figure 56,
Part B.

## Slide 145
a. Center gages check the angle of thread cutting tools, Figure 57.
The gage is also used to check cut threads and the scales are used
to measure the number of threads per inch.
b. Screw pitch gages measure the pitch of screws. If the pitch of a
screw is not known, it can be determined by comparing it with the
standards on various screw pitch gages as shown in Figure 58. Place
a blade of a gage over the threads, and check to see whether it
meshes; if not, successively check each blade of the gage against
the thread until it matches. The pitch can be read off the correct
blade. The blades are made pointed so they can be inserted in small
nuts to check inside threads as well as outside threads.

## Slide 148
Small Hole and Telescoping
a. Use these gages when measurements cannot be taken with a standard
micrometer, or when no plug gages are available.
b. Telescoping gages are particularly adaptable for rough bored work
and odd sizes and shapes of holes.
c. Use compress plungers and lock by turning the handle screw. Insert
the gage in the hole, Figure 59. Release the lock and the plunger
expands to the exact size of the hole. Lock the plunger by turning
the handle. After locking the plunger, remove the gage and check
the measurement with a micrometer.

## Slide 150
d. Small hole gages perform the same function as telescoping gages
except that they are used in small work.
e. To use the small hole gage, fit the ballshaped
point into the hole
or slot, Figure 60, Part A. Expand the ballshaped
end by turning
the screw at the end of the handle.
f. Use a micrometer to gage the measurement, Figure 60, Part B.

## Slide 151
Thread Cutting Tool
To check the correct form of a thread cutting tool, place the proper gage
over the tool, Figure 61. (The Acme standard gage is used here.) The tool
must mesh properly with no light showing between the tool and the gage. Use
a 29degree
angle as a guide when grinding a cutting tool.
When the tool fits the angle, grind the point to fit the proper place on the
gage. This indicates the number of threads per inch to be cut.

## Slide 153
Fillet and Radius
A double ended radius gage blade is used to check the inside corner or
fillet of a machine part, Figure 62, Part A. Each blade can be locked in
any position and at any angle for both inside and outside radii, Figure 62,
Part B.

## Slide 155
Drill Point (Rod)
Sharpen the cutting edges of a drill one lip at a time. Each lip must have
the same length and angle in relation to the axis of the drill. Set the
sliding head securely on the rule at the mark equal to the length of the
drill. Place the drill vertically against the rule so that the drill lip
contracts the 59 degrees angle of the sliding head. Hold the drill and gage
up to light. The correct angle is obtained when no light is seen between
the gage and drill.

## Slide 156
Lesson 5: TORQUE WRENCHES

## Slide 157
Learning Event
IDENTIFY TORQUE WRENCHES BY TYPES AND USE
A torque wrench is used for work requiring a particular force (torque). It
has a measuring device that indicates the amount of torque or twist being
applied to a nut or bolt. In much of today's machinery, engineers have
carefully worked out the degree of tightness to which nuts or bolts should
be tightened. If they are not tightened enough, they may loosen in service.
If they are tightened too much, the bolts or machine parts may be strained
excessively and break. In either case, the result could severely damage the
machine. The torque wrench allows nuts and bolts to be tightened just the
right amount

## Slide 158
Torque can be defined as a twisting force. A onepound
pull on a 1foot
wrench (center of bolt to point of pull) equals 1 footpound.
Force times
distance equals torque. The formula for determining torque would be as
follows:
(F x D = T) or 1 pound x 1 foot = 1 footpound
of torque

## Slide 159
When an extension (adapter) is used to lengthen a torque wrench, the
effective length is increased. Increasing the effective length of the
wrench changes the actual torque reading on the dial. To calculate the
actual torque, the following formula may be used:
A Torque
reading on dial
B Length
of torque wrench without adapter
C Length
of adapter

## Slide 162
TYPES OF TORQUE WRENCHES
Dial Indicating
The dialindicating
torque wrench, Figure 63, indicates the amount of torque
applied at any given moment. Dialindicating
wrenches are especially suited
for jobs requiring many different torque values because the dial does not
have to be reset each time. To measure bearing or preload, use a dialindicating
torque wrench.

## Slide 164
Deflecting Beam
Beamtype
torque wrenches, Figure 64, are made of flexible alloy steel which
deflects when torque is applied. A pointer attached to the end of the beam
indicates the torque reading on a scale. This type torque wrench has the
advantage of simple design. It does not use a multiplying device or a cam
mechanism that can be ruined easily

## Slide 166
Micrometer or Clicker
The micrometertype
torque wrench, Figure 65, has a builtin
micrometer at
the end of the handle where a specified torque can be set by turning the
handle. When the specified torque is reached, a loud click is heard. This
type torque wrench permits very accurate tightening of nuts and bolts

## Slide 168
Regardless of their type, most torque wrenches measure torque in footpounds
(ftlb).
However, some torque wrenches (mainly those used for more accurate
readings on smaller nuts and bolts) measure torque in inchpounds
(inlb).
Twelve inchpounds
equal 1 footpound.
To convert footpounds
to inchpounds,
multiply the footpound
by 12. To convert inchpounds
to footpounds,
divide the inchpounds
by 12.
Example of converting footpounds
to inchpounds:
75 footpounds
x 12 inchpounds
= 900 inchpounds.
Example of converting inchpounds
to footpounds:
1,260 inchpounds
¸ 12 inchpounds
= 105 footpounds.
Torque wrenches are issued in several sizes; they may have a 1/4,
1/2,
or
3/4inch
square drive to receive socket wrenches, and the capacity may range
up to 600 footpounds.

## Slide 169
SELECTING A TORQUE WRENCH FOR A SPECIFIC JOB
Before selecting a torque wrench for a job, note the torque specification.
Then select and use a torque wrench that will read about midrange for the
specified torque. For example, a typical 1/2inch
drive micrometer torque
wrench measures from 30 to 200 footpounds.
It would be the best torque
wrench for a job requiring about 100 footpounds
of torque. Typically, a
3/8inchdrive
torque wrench reads from 15 to 100 footpounds.
A 1/4inchdrive
torque wrench measures from 30 to 300 inchpounds

## Slide 170
USING THE TORQUE WRENCH
To use a torque wrench, set the specification on the micrometer handle or
note the specification on the scale or dial. Attach the proper socket to
the torque wrench. Place the socket on the nut or bolt to be tightened.
If an assembly containing many bolts or nuts is to be torqued, it is good
practice to use the following threestep
procedure:
1. First, torque all the bolts to about threequarters
of the
specification.
2. Recheck the specification and the torque wrench setting. Then,
apply the specified torque to each bolt.
3. Finally, retorque all bolts to the specification, to be sure that
none were missed in the final torquing.

## Slide 171
Use a smooth, steady pulling motion with a torque wrench. A fast or a jerky
motion will cause the fastener to be tightened to an incorrect torque. Be
sure to follow the torque specification for each fastener. Look for
footnotes, in the specification charts, that indicate whether the threads
are to be dry or lubricated with an antiseize compound. When you are
tightening parts such as cylinder heads or intake manifolds, always find and
follow the recommended bolt tightening pattern, Figure 66.

## Slide 173
CARE OF TORQUE WRENCHES
A torque wrench is a precision measuring instrument that accurately limits
the torque applied to a fastener. Do not use the torque wrench as a general
purpose turning tool. When using a torque wrench, observe the following
rules:
1. Never use a micrometer torque wrench in a position below the minimum
torque setting on the scale. However, when storing a micrometer
torque wrench, always adjust the handle to its lowest setting.
2. Never use the torque wrench to tighten a fastener to a higher torque
than the maximum reading on the scale

## Slide 174
3. Never use a torque wrench on a fastener that has been tightened with
a wrench or socket. For an accurate torque reading, the final
tightening must be done with the torque wrench.
4. A torque wrench has a recommended calibration interval. Most torque
wrenches are guaranteed to be accurate within four percent. To
ensure that the wrench is always this accurate, have it checked at
specified intervals by the manufacturer.
5. Examine and try torque wrenches. Note how varying amounts of pull
on the wrench handle register on the wrench dial. A torque wrench
should be pulled evenly and steadily; it should not be jerke

## Slide 175
6. Threads should be cleaned and lubricated before a nut or a bolt is
tightened. This reduces thread friction and results in more
accurate tightening.
7. Take special care of the torque wrench to avoid damaging it or
ruining its adjustment. Keep it clean and oiled; put it away in a
protected place after each use.

## Slide 176
METRIC SYSTEM VERSUS US CUSTOMARY SYSTEM
Today, knowing how to convert measurements from the US customary system to
the metric system is very important.
Kilograms (kiloponds) or newtons express force in the metric syste

## Slide 177
Use the appropriate type of torque wrench for specifications given. If
specifications are given in centimeterskilograms,
then a torque wrench with
a centimeterkilogram
scale should be used. If torque values stated do not
match the valves on your torque wrench, the following formulas can be used
to convert the values given:

## Slide 179
NOTE: The newtonmeter
is a measurement of torque used by any countries
using the metric system. One newtonmeter
is equal to .7375 footpounds.
To convert newtonmeters
to footpounds,
you multiply the total newtonmeters
by .7375.
Example: 150 newtonmeters
x .7375 = 110.6 footpounds
To convert footpounds
to newtonmeters,
you would divide the total footpounds
by .7375.
Example: 80 footpounds
¸ .7375 = 108.5 newtonmete

## Slide 180
Lesson 6: HYDRAULIC FILLER/BLEEDER

## Slide 181
USING THE HYDRAULIC FILLER/BLEEDER
The hydraulic filler/bleeder is used to purge (remove trapped air) a
hydraulic system such as a hydraulic brake or a hydraulic clutch system.
When a hydraulic brake or clutch system must have a component repaired or
replaced, the system must be opened. This allows air to enter the system.
Air in the system results in spongy pedal feel and action. Since air in the
system is compressible, reduced pressure and force result in poor braking
action when the brakes are applied.
Although you can bleed a hydraulic brake or clutch system without a
filler/bleeder, using it simplifies the job, and allows you to accomplish
the task without assistance.

## Slide 182
The hydraulic filler/bleeder can be used to bleed hydraulic brake systems,
hydraulic clutch systems, dual safety cylinders, dual clutch and brake
cylinders, all power brake systems, and any master cylinder with or without
a check valve.
DESCRIPTION
The filler/bleeder, Figure 67, is a round cylinder divided into two chambers
by a tank diaphragm. The top chamber holds hydraulic brake fluid. The
bottom holds pressurized air. The filler/bleeder is mounted on casters
(rollers) which allow it to be moved easily around the shop.

## Slide 184
The filler/bleeder comes with seven bleeder adapters, Figure 68. They are
designed to fit most master cylinders on the market at this time.

## Slide 186
In addition to the adapters, the filler/bleeder comes with bleeder hoses,
elbows, male and female quickdisconnects,
and a spare gasket set, Figure
69.

## Slide 188
The numbers that appear on the filler/bleeder and its adapters are parts
numbers which correspond to the parts list in Figure 70.

## Slide 190
Filler Unit
a. Refer to Figure 67. Loosen valve cap A to exhaust the air chamber.
Leave valve A open. Also, ensure that valve B is fully open.
b. Wipe dirt from the top of the tank and open filler plug C, Figure
67. Fill the fluid chamber to the brim of the inlet with brake
fluid (silicone, automotive all weather, operation and preservative,
military specification MilB46176).
c. Close plug C, Figure 67, hand tight. Tighten air cap A securely and
fill the unit with air through cap A until the gage reads in the
black, 20 to 30 psi (20 psi is sufficient to bleed all vehicles).

## Slide 191
WARNING
DO NOT CHARGE THE UNIT WITH AIR IN EXCESS OF 50 PSI.
DO NOT USE SHARP TOOLS OR OTHER PROBES TO RETURN
DIAPHRAGM TO ITS NORMAL DOWN POSITION.
d. Open bleeder screw D, Figure 67, to bleed any air that may be
trapped in the fluid section of the filler/bleeder.
e. On the initial filling only, the hose should be bled using the male
quickdisconnect
fitting, part number T1712, Figure 69.

## Slide 192
Connecting Filler/Bleeder to Master Cylinder
a. Clean the top of the master cylinder to keep dirt from dropping into
the opening. Remove the filler cap from the master cylinder and
clean the gasket seat thoroughly.
b. Refer to Figure 68. Select and install the proper adapter and
gasket. On most adapters, hand tightening is sufficient with clean
surfaces and the proper gaskets. Others may require a wrench to
tighten enough to hold the proper pressure.
c. Screw the quickdisconnect
male fitting, part number T1712, Figure
69, into the adapter on the master cylinder and hand tighten

## Slide 193
d. Connect the female quickdisconnect,
Figure 71 (the female
disconnect is attached to the end of a long hose), directly to the
male adapter on the master cylinder.
e. Be sure valve B, Figure 67, is open. The system is now ready for
bleeding

## Slide 195
General Procedures
a. Check for leaks at all line, hose, and cylinder connections you made
on the vehicle.
b. Bleed units in the following order: master cylinders (some have
bleeder screws), booster or power units, line bleeder screws, and,
finally, all wheel cylinders.
c. Refer to Figure 72. Attach the bleeder hose, part number B75 or
B5156, to each bleeder screw in turn and insert the open end of the
hose into the bottom of a glass container. The end of the hose
should be submerged in the fluid. Open the screw to permit all air
to escape. Watch for air bubbles in fluid for a positive check.
When fluid flows clear without air bubbles, close the bleeder screw
and tighten it securely.

## Slide 197
d. To remove most of the air from the master cylinder quickly, you
should bleed the left front wheel cylinder first, then the right
rear, the left rear, the right front, and the left front again. It
is good practice to rebleed the four wheels again to ensure a
perfectly airfree
system.
e. To disconnect the bleeder, always cover the quickdisconnects
with
cloth before disconnecting them. This avoids fluid spilling onto
the vehicle finish. Just uncouple the quickdisconnect,
and remove
the adapter from the cylinder. Be sure the reservoir is filled to
within 1/2 inch or less of the top and replace the gasket and filler
cap.

## Slide 198
Caution
Always cover fenders and keep hoses clean to avoid
damaging finish.
Power Brakes
a. Follow the general procedures above, except DO NOT bleed with the
motor running.
b. If the power unit has a bleeder screw, bleed it first. Some have
two bleeder screws. When there are two, bleed at the highest screw
first.

## Slide 199
c. When bleeding the treadle valve, use only low pressure, 10 to 12
psi, in the bleeder tank.
d. When bleeding the hydrovac
double stage control valve, rebleed the
highest bleeder screw three or four times to ensure complete air
removal.
Clutch Systems
a. Follow general bleeding procedures.
b. Clutch controls will always lose pressure and will not build up like
brake pedals do since pumping will not help.

## Slide 200
Vertical Type Wheel Cylinders
a. Cylinders with an internal cross connecting line should be prefilled
before installing. Connect cylinders with the crossline
hand
tightened on the bench and prefill the cylinders. Moving the
pistons slightly will help to assure complete filling.
b. Hoses and lines should be pressure bled before connecting the wheel
cylinders. Rebleed units after installing.

## Slide 201
Surge Bleeding
Some vehicles have diagonal or vertical wheel cylinders which require surge
bleeding. On such cylinders, the bleeder screw is not always the highest
point to allow air to escape, therefore, complete bleeding is difficult.
Surge bleeding will help:
a. Bleed each wheel with the pressure bleeder.
b. Open one bleeder screw at a time and have someone kick the brake
pedal vigorously a few times. This will cause turbulence in the
cylinder forcing additional air out.

## Slide 202
Caution
Before allowing the vehicle or equipment to be
operated, check the following:
Always check connections for leaks by applying heavy
pressure on the pedal or treadle. If the system is
vacuum or air assisted, start the engine to create a
vacuum, or to build air pressure.

## Slide 203
Always check the bypass
part in the master cylinder
to be sure it is clear when the system is at rest.
Never probe in the small bypass
hole. When applying
the brakes, there should be a ripple in the fluid in
the master cylinder reservoir. After applying brakes
a few times, there should also be a ripple caused by
returning fluid. Checking visually with a bright
light is advisable.
WARNING
IF THE BYPASS
PART IS NOT CLEAR, BRAKE SYSTEM LOCK UP
COULD OCCUR AND CAUSE SERIOUS DAMAGE.
Caution
Always check the brake pedal or linkage for free
travel. Adjust in accordance with the appropriate
technical manual as needed.

## Slide 204
FILLER/BLEEDER MAINTENANCE
Bleeder/Filler Tank
The tank and fittings should be kept clean at all times. Never allow the
tank or fittings to come into contact with mineral oil products such as
kerosene, motor oil, and thinners. Do not use water to clean the tank at
anytime.
Fittings
Fittings cleaned in anything other than brake fluid should be thoroughly
dried inside and out before use.

## Slide 205
Air Chamber
Periodically clean out the air chamber in order to remove any water, oil,
and dirt sludge that may enter from the air system used to fill the tank.
Remove and rinse the safety valve (part number T4669) with clean alcohol.
Be sure to drain and dry it completely to avoid contaminating the next fluid
fill.

## Slide 206
Fluid Chamber
The fluid chamber needs no attention unless dirt or improper fluid has to be
removed. Rinse it with clean alcohol, and drain and dry the chamber
completely to avoid contaminating the next fluid fill.

## Slide 207
Safety Valve
The safety valve (part number T4669), Figure 70, on the bottom of the tank
is set to blow off at about 50 psi. When pressure is relieved, it will
reseat itself at about 30 psi.